- Principle

Maintenance will be performed daily, weekly and monthly as per Immucor protocol.

- Scope and Related Policies – N/A

- Specimen – N/A

- Material

Equipment:	Immucor Galileo NEO

Supplies:	De-ionized water

70% Isopropanol

Virkon

Lint free cloths

Plate carriers

Balance

Calculator

Reagents:	PHIX buffered Saline

Capture-R Ready Screen (3)

- Quality Control

Review of the NEO maintenance log is performed weekly by the administrative user and monthly by the MLT 3.

- Procedure

  - Daily Maintenance

    - Refill PBS supply container.

    - Empty liquid waste container.

    - Perform a NEO module and PC shutdown by pressing exit door icon on main NEO screen.  Click on “Shutdown”.  Power off NEO instrument.

    - During brief shutdown clean the instrument including the pipettors using 70% isopropyl alcohol and wipes.

    - Immediately power back on PC.  Wait for main screen and power on NEO.

    - Wait for initialization to complete.

    - Click on maintenance menu.

    - Select Clean Instrument from the list. Press Start and Continue on the following screen.

    - Select “Check pipettor reference”.  Click start.

    - Press the Check Reference button.

    - Check both Left and Right pipettor arms.

    - Probe tip should fall within the inner circle.

    - If the alignment is OK.  Press Position OK button and proceed with other arm.

    - If alignment is off.  Press Position Wrong.  Make necessary adjustments to probes and restart verification.

    - When done select the close button.

    - Select Verify Reader from the list.

    - Press the Start button.  Load the RVP plate into the plate tower.

    - Under plates manually assign a plate ID and choose RVerify. Press Start.

    - When done return RVP to its case.

    - Select Pipettor Self Check from maintenance list.

    - Load clean tool into plate tower.  Close the door and let the instrument identify plate.

    - Press Start.

    - Clean tool after completion.

    - All daily maintenance must be updated to OK for assays to be run. If maintenance is Not Passed or Due assays will not run.

  - Weekly Maintenance

      - Verifying the Washer:

        - Press the Maintenance button.

        - Select Verify Washer.

        - Press Start.

        - Follow the on screen instructions.

        - Flow350 and Flow200 must be run.

        - Load the washer Flow Verification Tool.

        - Under plates manually assign a plate ID and choose Flow350.

        - Press Start.

        - Once complete verify that the volume is between 2.0 an3.0 ml.

        - Follow same procedure for Flow200.

        - Clean Tool and return to Neo cabinet

      - Reader Performance:

        - Press the Maintenance button.

        - Select Reader Performance.

        - Press Start.

        - Press the Take flatfield button.

        - Verify the last screen image for any significant spills. Bring to the attention of the Key operators. The mirror or light diffuser may need to be cleaned.

        - Click on done.

        - Load RVP plate. Manually assign a plate ID. Select RPerform. Press Start

        - Once completed return the RVP plate to its case.

      - Archiving:

        - This task must be performed on a weekly basis or the system may become unstable.

        - Click on File Management on the Tool Bar.

        - Select utilities on the menu bar.

        - Insert Blank DVD into drive.

        - Verify enough space is available on the disk.

        - Select the GO Button.

        - Wait for the Archive complete box to appear. Click OK.

        - Label disk.

      - Cleaning the Pipettor Wash Towers:

        - Prepare a 1 L working solution of Virkon. (1% 2 tablets/L)

        - Shut off the NEO.  Manually move all pipettors to the centre of the instrument.

        - Dispense the solution into both wash tower rinse cups.

        - Expose for 10 minutes

        - Clean common waste Container

      - Cleaning the Common Waste Container

        - Prepare a 1L 1% suspension of Virkon.(Use remaining solution from 6.2.1.4)

        - Empty the Common Waste Container

        - Rinse the container with the Virkon solution.

        - Expose for 10 minutes.

        - Clean the inside of the container lid with isopropanol.

        - Rinse the container with deionised water.

        - Empty the container.

        - Reconnect the container to the instrument.

        - While the NEO is turned off, inspect the syringes for tightness and leakage.

        - Power on the NEO and allow the lamp to warm up for a minimum of 30 minutes.

  - Monthly Maintenance

    - Check the RVP Calibration date.  This can be found on the RVP

plate case

    - Add 3 years to the calibration date and note on the maintenance

sheets.

    - Inform Technical Support if calculated date is within one month of

Expiry

    - Perform Decontaminate Tubings procedure as per NEO Operator

Manual Monthly Maintenance section.

    - The Monthly Decontamination can take ≥ 1 hour to complete and is

usually performed by the NEO Key Operator.

  - As needed Maintenance

Shall be performed by Immucor Technical Services or NEO key operator on an as needed basis only.

- Reporting

  - Daily, weekly and monthly maintenance sheets are kept on the NEO bench.

- Procedural Notes

  - Daily and weekly maintenance are timed and must be performed.  The instrument will not perform any testing if maintenance times have lapsed.

- References

  - Galileo NEO Operator Manual