- Principle

Anti-Ch and Anti-Rg are antibodies directed against the C4 component of human complement. Neutralisation of these antibodies can be achieved by incubation with plasma from antigen positive individuals.

- Scope and Related Policies

  - Chido and Rodgers antigens are epitopes on the C4 component of human complement.

  - Plasma that contains C4 from Ch/Rg postivie individuals can be used to inhibit the reaction of Ch/Rg antibodies.

  - Plasma that has undergone inhibition can be used to identify the presence of other alloantibodies that may have been previously masked by the Ch/Rg antibody.

- Specimen

EDTA anticoagulated whole blood preferably less than 72 hours old.

- Material

Equipment:		Cell Washer

Serological centrifuge

Block for test tubes

Microscope

Waterbath/Heating block at 37°C

Supplies:		Test tubes – 10 x 75 mm

Serological pipettes

Reagents:		0.9% saline

Anti-IgG

IgG-coated cells

Inert Plasma – Group compatible (Pool of at leastsix normal plasma samples known to be non-reactive with screening cells).

Test cells (antibody screen cells or a panel of cells for antibody identification).

6% Bovine serum albumin (BSA).

- Quality Control

  - A control consisting of the indicator red cell and the pooled neutralising plasma must always be included.

  - The absence of activity in the 6% albumin control indicates dilution of a weakly reactive antibody and an invalid test.

- Procedure

- Prepare serial twofold dilutions of patient’s plasma in saline. The dilution range should be from 1 in 2 to 1 in 512, or to one tube beyond the known titer. Refer to SP.002 – Titrations. The volume prepared should be not less than 0.3 mL for each red cell sample to be tested.

- For each red cell sample to be tested, place 2 drops of each plasma dilution into each of two sets of appropriately labeled 10 x 75 mm test tubes.

- To one set, add 2 drops of inert pooled plasma to each tube.

- To the other set, add 2 drops of 6% albumin to each tube.

- Gently agitate the contents of each tube and incubate the tubes at room temperature for at least 30 minutes.

- Add 1 drop of a 3% suspension of red cells to each tube.

- Gently agitate the contents of each tube and incubate the tubes at 37°C for 1 hour.

- Wash the cells four times in saline, add anti-IgG, and centrifuge for 10-15 seconds at 3400 rpm.

- Resuspend the cell buttons and examine for agglutination; confirm all non-reactive tests microscopically. Grade and record the results.

- Confirm the validity of negative results by adding IgG-coated red cells.

- Reporting

  - Inhibition of antibody activity in the tubes to which plasma has been added suggests anti-Ch or anti-Rg specificity; this inhibition is often complete.

  - The presence of partial inhibition suggests the possibility of additional alloantibodies. These antibodies should be identified before reporting the final results, see table below:

Interpretation of Results:

| TEST PLASMA | CONTROL PLASMA | INTERPRETATION |
| --- | --- | --- |
| No reactivity with all cells tested | Reactivity with all cells tested | Antibody neutralised specificity Chido or Rodgers |
| No reactivity | No reactivity | Dilution of antibody has occurred No conclusion can be made |
| Reactivity with some cells tested | Reactivity with all cells tested | Anti-Ch or Anti-Rg plus an additional antibody maybe present |

  - To confirm the presence of additional antibody(ies) prepare a large volume of inhibited plasma and test it against a reagent red cell panel to determine whether the non-neutralized activity displays specificity towards a particular antigen.

- Procedural Notes

  - Some red cell related HLA antibodies (e.g. Bg) may be neutralised with human plasma.

  - Plasma may also inhibit anti-Cromer.

- References

  - Judd, WJ ed. Judd’s Methods in Immunohematology, 3rd ed,   Bethesda, MD: pg. 345-347.

  - Roback, JD. ed. AABB Technical Manual, 17th ed. Bethesda, MD: American Association of Blood Banks, 2011: pg 911-912.

- Revision History

| Revision Date | Summary of Revision |
| --- | --- |
| September 1, 2014 | - Revised name of manual - Revised section 2.0 - Replaced “red cells” with “whole blood” in section 3.0 - Replaced “Normal Saline” with “0.9% Saline” in section 4.0- Reagents - Revised wording of section 6.1; added section 6.10 - Updated list of references to include most recent editions |